Ammonium (NH4+) Protocol

(Modified from Strickland and Parsons)

The following method is specific for ammonia. The technique is a modification of the phenol-hypochlorite method published by Solarzano, L. (1969).

# Introduction

Nitrogen, in the form of NH3 (ammonia gas) or NH4+ (ammonium ion), is the first breakdown product of N-containing organic matter. A sensitive colorimetric test for ammonia nitrogen in seawater was not developed until 1969. Prior to that time samples had to be painstakingly distilled before analysis. The method presented here is suitable for both seawater and freshwater although for highest precision at low concentrations some modifications may be made for the analysis of freshwater.

The method employs the so-called ‘indophenol-blue reaction between ammonia, phenol and hypochlorite in an alkaline medium.  Indophenol is an intensely colored amino-derivative of phenylquinone-monoamine:

The resulting blue color is similar to that of the indophenol dyes and its intensity is proportional to the ammonium concentration. The indophenol colorometric reaction was modified by the introduction of the catalyst nitroprusside which accentuates the blue color at room temperature.

NOTE: Due to the ubiquity of ammonia gas, test tubes are easily contaminated during long-term storage. Test tubes should be pre-reacted (see next page) one or two days before the lab and allowed to dry.

Solutions:

- Phenol solution:

In a fume hood, dissolve 20 g of crystalline analytical reagent grade phenol in 100 mls of 95% v/v ethyl alcohol (or 22 mls 90% Phenol in 200 ml 95% ETOH). This solution is stable for several months.

- Sodium Nitroprusside Solution:

Re-crystallize sodium nitroprusside: dissolve 5g of sodium nitroprusside in 10 ml of DI water. (This can take several hours). All the nitroprusside may not dissolve. Use a stir bar to help dissolve crystals. Add some DI if necessary to help dissolve. Add 25 mls of 95% ethyl alcohol.  Freeze overnight to reform crystals and then filter (use cold ETOH to wash crystals) and dry the crystals in a dessicator.  Re-crystallized sodium nitroprusside will give lower reagent blanks. Dissolve 1.0 g of re-crystallized sodium nitroprusside in 200 ml of DI water.  Store in an amber bottle.  The solution is stable for at least one month.

- Alkaline Reagent:

Dissolved 100 g of sodium citrate and 5 g of sodium hydroxide (analytical reagent grade) in 500 ml of DI water.  This solution is stable indefinitely. Use caution when handling NaOH.

- Standard Ammonia Solution (10,000 M stock):

Dry 5 grams of NH4Cl in an oven set at 60 C for 24 hours.  Once this is cool, add 0.5349 g of dried NH4Cl to a 1 L volumetric flask and fill to the line with DI. Add 1ml chloroform to preserve. This solution is 0.01 M NH4Cl (10,000 uM NH4Cl).  This solution is stable for many months if well stoppered.

Each Day before samples are run:

Make up Oxidizing Solution:

Mix 100 ml of the Alkaline Reagent with 25 ml of Sodium Hypochlorite Solution (chlorine bleach).  Keep solution stoppered while not in use.  Prepare fresh every day.

Make up a set of standards using the 10,000 M NH4Cl stock solution.  Include a DI blank with your standard curve.  We often use the concentrations of 0, 0.5 M, 1 M, 5 M, 10 M, 50 M, and 100 M for our measurements.  The analysis is not linear beyond 100 M.

Turn on the spectrophotometer and allow warming up for 20min.  Set the wavelength to 640nm.

If you suspect you have high concentrations of ammonia in your samples, dilute as necessary to bring them within range of your standard curve.

IN THE HOOD:

To 5 ml of sample add the following reagents (using a repipettor makes this go quickly).  Mix the test tube using the vortex mixer after each reagent is added.

- 0.2 ml of Phenol Solution,

- 0.2 ml of Sodium Nitroprusside Solution,

- 0.5 ml of Oxidizing Solution.

Note: Other sample volumes can be used (e.g. 3 ml) and reagent volumes are then adjusted proportionately (e.g. for 3 ml sample use 0.12 ml Phenol and Nitroprusside solutions and 0.3 ml Oxidizing solution).

Cover tubes with parafilm.  Allow the samples to stand at a temperature between 20 C and 27 C for at least 1 hour in the dark (samples should be analyzed within about 12 hrs, for example samples can be left to develop overnight).  Color will continue to intensify over several hours, so standards and unknowns should be measured in approximately the same time frame.

Read the absorption at 640 nm on spectrophotometer.

NOTE Regarding contamination

For low level ammonia analysis, only use test tubes that have been pre-reacted with the reagents to remove any sorbed ammonia.  Plastic 15 ml conical centrifuge tubes may be less subject to contamination and preferable to use for ammonia determinations.  To pre-react new tubes: rinse tubes three times with DI water. Pipette 5 ml of DI water into each test tube.  IN THE HOOD: add reagents: 0.20 ml of phenol, 0.20 ml of sodium nitroprusside solution, and 0.50 ml of oxidizing solution to each test tube, using the vortex to mix after each reagent is added.  Allow pre-reacting tubes to stand at room temperature for 1 hour.  Invert tubes (make sure you are wearing gloves) to allow reagents to contact all surfaces of the tubes.  Dispose of waste properly.  Fill tubes once with DI water and empty into the NH3 waste container in the hood.  Rinse tubes twice more with DI water (empty tubes in the sink).

Allow tubes to dry completely, and cover with parafilm to avoid contamination by atmospheric NH3.

Dump waste into clearly marked container labeled: NH4+ waste: phenol, ethanol, sodium-Nitroprusside, sodium-Citrate, NaOH

Clean up!!!

References:

Solarzano, L. (1969). Determination of ammonium in natural waters by phenol hypochlorite method. Limnol. Oceanogr., 14:799-800.

Strickland,J.D.H. and T.R.Parsons A practical handbook of Seawater Analysis 1972 Ottawa, Fisheries Research Board of Canada 2nd. Ed.

Reprinted From Limnology and Oceanography (1969) Vol. 14:799-800