Lachat Setup – from Kelsey Chenoweth

- Make sure column is off

- Turn on power strip in back

- Tubes in DI -> click down cassettes and pull levers up. Check flow. Adjust levers if needed.

- Manual on/off @ 35, Fresh DI in carrier flask (unless using salt water as carrier)

- Turn on computer (pw=lachat) and open Omnion

- Highlight all but standard curve and checks -> right click -> delete

- Right click -> insert many -> highlight all but checks -> right click ->columns -> autonumber cups

- Rename by highlighting -> right click -> auto sample ID

- Copy first check -> insert every 20 samples

- Save as in correct location (may need to exit and reopen)

- Configuration -> options -> data export -> data items -> save to correct location

- Run properties -> sample -> click on standards to check the value. Racks -> check for proper size.

- Configuration -> Autosamplers -> initialize -> prime

- Turn off pump -> transfer lines to reagents -> parafilm

- Wait ~2 minutes, check baseline, then turn column on

- Pour standards, Nitrite check, DI and place (1,2,3,4,5 in back, 6,7, 8,9,10 front)

- Check that sample tray is clicked in place

- Preview (eye symbol) -> check that it’s around 1 -> Stop -> Start

Shutting down the Lachat

- Turn off column

- Stop pump -> tubes in DI -> Start pump

- Check for file in folder (open excel to see data)

- Save both .omn and .csv to USB drive

- Close Omnion and turn off computer

- Tubes lifted out of the DI, run air through (column is off!). Turn autosampler off and on to pump air.

- Turn off Lachat with power strip

- Release levers and lift cassettes!!